Hybrid Un-Tailing & Un-Plating Certification

( The following individuals can train others in the proper procedures

for hybrid untailing and un-plaitng:

Derek Barge

( The following individuals have been trained in the hybrid handling

procedures by someone authorized to do so in the list above.

| Trainee | Trainer Autherization | Date |
| --- | --- | --- |

# I. General Cleanliness

1) Hands:

- Gloves or finger cots must be used whenever a hybrid is

touched. DO NOT touch any part of any hybrid with a part

of your hand that is not protected by gloves or finger cots.

2) Face & Hair:

- A face mask must be worn at all times while handling hybrids,

and at all times working near an exposed hybrid (an exposed

hybrid is one that is not in a sealed hybrid box). The face mask

must fully cover the mouth and nose.

- A hairnet must be worn inside the clean rooms at all times.

3) Feet:

- Clean booties must be worn inside the clean rooms at all times.

- Grounding straps must be worn any time a hybrid is being

handled. Contact must be checked at a DESCO ‘touch tester’

or at the large DESCO ‘WRIST STRAP / FOOTWEAR tester’ in

the main clean room before handling hybrids.

3) Body:

- A lab coat must be worn inside the clean rooms at all times.

Make sure your lab coat was designed for clean room use.

# Hybrid Untailing & Un-plating Procedure

- If anything does not go according to this procedure, or you are

unsure about what you are doing, stop immediately and contact a

supervisor. DO NOT continute to work until a supervisor is available.

*This procedure should be completed without touching the hybrid

except by its tail (using gloves).

*No tools, fingers, or other objects should be held directly above the

hybrid at any time.

* This procedure should only be completed for hybrids that have successfully passed

the tests at the 4-hybrid test stand.

1) Check to make sure you are wearing the following:

- booties

- gloves

- hairnet

- lab coat

- face mask

- grounding straps

2) Make sure the work surface is clean and free of clutter. Remove the plated hybrid

from the hybrid box and set on the work surface.

3) Carefully remove the bottom strap.

4) Cautiously, remove the hybrid tail. Follow a) for hybrids with up connectors and b)

for hybrids with down connectors. Use a flat head jewlers screw driver with a 1/8"

blade.

a) Hybrids with ‘up’ connectors

- Take a small flat head jewelers screw driver and place the edge

of the flat head in the top seam between the hybrid connector and

the test tail. Make sure not to pry the connector mounted

on the hybrid tail against the hybrid tail.

- Rotate the driver 90(, this will free one side of the connector.

- Very gently push down on the SAME SIDE side of the connector

that was just freed in the upper right hand corner, this should release the test tail

from the connector.

- If the test tail does not come completely free, gently grasp both ends and

gently wiggle it free. There should be very little to no resistance.

- Gently remove the hybrid tail.

b) Hybrids with ‘down’ connectors

- Take a small flat head jewelers screw driver and place the edge

of the flat head in the bottom seam between the hybrid connector and

the test tail. Make sure not to pry the connector mounted

on the hybrid tail against the hybrid tail.

- Rotate the driver 90(, this will free one side of the connector.

- Very gently push down on the SAME SIDE side of the connector

that was just freed in the upper right hand corner, this should release the

connector from the test tail.

- If the hybrid tail does not come completely free, gently grasp both ends and

gently wiggle it free. There should be very little to no resistance.

- Gently remove the hybrid tail from the test tail and remove test tail.

4) If the hybrid has a ‘Down’ connector proceed to step 5. If the hybrid has an ‘Up’

connector, carefully remove the upper strap.

- Extra caution should be used because of the close proximity

to the hybrid.

- After unthreading each screw, but before removing it, make

sure to hold the screw in place while seperating the screw from

the driver. This is to prevent the screw from sticking to the

driver and then falling on or near the hybrid. Loose allen

screws can damage hybrids.

- Remove the left screw followed by the right. This is so that the

strap cannot rotate towards the hybrid and damage it.

5) Loosen the allen screws securing the hybrid clamps on either side of

the hybrid (1/2 turn or so should be sufficient).

- DO NOT completely remove the allen screws from the plate,

loose allen screws could cause damage to the hybrid.

6) Carefully slide the hybrid away from the plate and onto the work

surface.

7) Return all screws and clamps to their designated bins. Return the plate to the

hybrid box.

8) Gently grasp the hybrid by the tail and place in the center of a small

hybrid box.

9) Form a loop at one end of a 4” piece of kapton tape. This is for easy

removal later.

10) Without touching the hybrid, apply one piece of kapton tape across

the tail and one corner of the hybrid.

11) Close the hybrid box. Make sure to transfer any labels from the original box to

the new box. If there is a serial number labeled on the outside of the box

remove it. Now Place on the correct shelf in the cabinet containing tested

hybrids. The hybrid is now ready for gantry assembly.