OGP survey Certification

( The following individuals can train others in the proper procedures

for surveying modules on the OGP.

Dean White

Andrea Allen

( The following individuals have been trained in the OGP survey

procedures by someone authorized to do so in the list above.

| Trainee | Trainer Authorization | Date |
| --- | --- | --- |

# After cure survey on the OGP

- Wear facemask, hair net, lab coat, and grounding straps.

- Drive out OGP stage.

- Place assembly plate onto the OGP stage, push all the way against the pins, and clamp the assembly plate into place.

- Open MeasureMind Plus and SmartReport Plus.

- [In SmartReport, click on the export tab. Configure export should start at “Actual” cell C:8 . The data should also be exported in rows, not columns. This only needs to be done if you are opening SmartReport fresh, if you have previously ran the program successfully you should not have to do this]

- Open the correct MeasureMind program, it should be the program corresponding to the particular plate being used where the file is named  plate##-smrt.rtn  and ## is the assembly plate number. Make sure  you know what type of sensors are being used, there should be a different program for ST and HPK sensors.

- Click on the circle tool, you will center the fiducials in this circle when initially zeroing the axes. For best results zoom in all the way.

- Press run, it’s on the upper left of the screen it looks like a play button.

- Find the fiducial on the left side of the assembly plate and center it in the circle target tool, or any other target tool.

- Zero the x, y, and z axes by either clicking the x, y, z buttons on the bottom of the screen or by pressing the buttons on the OGP control panel.

- Align the x axis through the fiducial on the right side of the assembly plate by pressing ANG on he screen or pressing the axis align button on the OGP control panel. These last two steps are important, if they are not done before a run it could cause a crash of the OGP into the assembly plate.

- Let the program run, watch while it is running to make sure it catches all the fiducials, if not make a note of the ones it misses to back and re-measure manually.

- When the program is finished running minimize the MeasureMind window. Smart Report will working the background. Excel will then pop up, run the macro by going to the tools menu select Macros and then run Main.

- Once the macro is running you will be asked to enter your name, the temperature, the humidity, and then scan the module numbers in each position.

- The files will print out and be saved in a text file that gets uploaded in to the database.