TOB Module Tailing Certification

( The following individuals can train others in the proper procedures

for TOB module tailing.

Derek Barge

( The following individuals have been trained in the hybrid handling

procedures by someone authorized to do so in the list above.

| Trainee | Trainer Authorization | Date |
| --- | --- | --- |

# I. General Cleanliness

1) Hands:

- Gloves or finger cots must be used whenever a hybrid is

touched. DO NOT touch any part of any hybrid with a part

of your hand that is not protected by gloves or finger cots.

2) Face & Hair:

- A face mask must be worn at all times while handling hybrids,

and at all times working near an exposed hybrid (an exposed

hybrid is one that is not in a sealed hybrid box). The face mask

must fully cover the mouth and nose.

- A hairnet must be worn inside the clean rooms at all times.

3) Feet:

- Clean booties must be worn inside the clean rooms at all times.

- Grounding straps must be worn any time a hybrid is being

handled. Contact must be checked at a DESCO ‘touch tester’

or at the large DESCO ‘WRIST STRAP / FOOTWEAR tester’ in

the main clean room before handling hybrids.

4) Body:

- A lab coat must be worn inside the clean rooms at all times.

Make sure your lab coat was designed for clean room use.

# TOB Module Tailing Procedure

- If anything does not go according to this procedure, or you are

unsure about what you are doing, stop immediately and contact a

supervisor. DO NOT continute to work until a supervisor is available.

*This procedure should be completed after module bonding and before ARCS testing.

*No tools, fingers, or other objects should be held directly above the

module at any time.

1) Place the module on a blue ESD mat.

2) Tape down the high voltage flying wire.

3) If the module has an ‘Up’ connector, carefully secure the upper low voltage clamp.

Otherwise proceed to step 4.

- Example of module with ‘Up’ connector:

- Example of module with ‘Down’ connector:

4) Carefully secure the upper high voltage clamp.

- Example of module with ‘Up’ connector:

- Example of module with ‘Down’ connector:

5) Attach the high voltage tail.

6) Attache the low voltage tail.

- Example of module with ‘Up’ connector:

7) If the high voltage is on the right, secure the bottom of the high voltage and low

voltage tails with a single ‘TOB WIDE’ clamp:

- Example of module with high voltage on the right and ‘Up’ connector:

- Example of module with high voltage on the right and ‘Down’ connector:

Otherwise the high voltage is on the left and the lower high voltage and low voltage tails must be clamped seperately:

- Example of module with high voltage on the left and ‘Up’ connector:

- Example of module with high voltage on the left and ‘Down’ connector:

8) The module is now ready for ARCS testing. Place the module on the proper shelf

in the cabinet labeled ‘ARCS TESTING MODULES’.